HTX9000 CALIBRATION PROCEDURE

John Cook    April 2013

Purpose:  This document describes how to calibrate the HTX9000 automatic USB connected load box and how to read back the calibration data.

Items needed:

HTX9000 USB controlled load box

USB cable with “mini-B” connector.

Power supply able to produce at least 2.5 amps of current at least 5 VDC

Accurate current meter with Amps and milliamp ranges

Jumper cables.

Procedure:

- Connect the load box to a PC via USB cable.  It should power up, show its logo, and show the load setting.  If the load box is uncalibrated, ignore the load setting on the display.

- On the PC, determine the virtual COM port attached to the load box and open a terminal program on that COM port to communicate with the device.

- In the terminal program, send the command “*IDN?”  The load box should respond with its make and model information.

- Set the power supply to 5VDC, connect it to the load box with the current meter between them.

- Send the following commands to the load box, in the following order:

  - CALibrate:DATE (today’s date, ex: 4/26/13)

  - CALibrate:OUTPut:CURRent:DAC:RANGe:HI 65535

  - CALibrate:OUTPut:CURRent:VALue:RANGe:HI  (meter reading, in amps.  Ex: 2.4998A)

  - CALibrate:OUTPut:CURRent:DAC:RANGe:LO 65535

  - CALibrate:OUTPut:CURRent:VALue:RANGe:LO (meter reading, in mA. Ex: 24.99962mA)

- While calibrating, the display will be locked to prevent changes on the load box’s control buttons.

- After calibration, command the load box to various load currents with the following command: “SOURce:CURRent 1A” or “SOURce:CURRent 10mA” and observe the meter measuring the commanded current.

- The Board can be given a version number or title limited to 16 characters such as “Steve’s LoadBox” with:

  - CALibrate:BOArdrev Steve’s LoadBox

- Mark the meter as calibrated with a dated calibration label.